Lab::VISA
Lab::VISA
& friends
slides © Daniel Schröer and Andreas K. Hüttel

---
VISA, GPIB, etc.
VISA, GPIB, etc.
Instruments can be connected in various ways:
Serial port, GPIB, VXI, TCP/IP, etc.
GPIB (hardware and software)
  - GPIB (IEEE488): Standard by Hewlett-Packard
  - Physical layer IEEE488.1
  - Command layer IEEE488.2
  - SCPI (Standard Commands for Programmable Instruments)
VISA (software)
  - Virtual Instrument Software Architecture
  - VXI, GPIB, serial, or computer-based instruments
  - NI-VISA library is one implementation of the VISA standard
[Speaker Notes: most important for us: GPIB]

---
Lab::VISA
- Lab::VISA is software to control instruments in the lab via VISA
  (e.g. voltage sources, multimeters, etc.)
- Alternative to LabView, Gpplus, etc.
- But very different: no GUI, just API

---
Lab::VISA design goals
Lab::VISA design goals
Flexible
  - Allow any kind of measurement procedure
  - Control anything that has GPIB or serial connection
Safe
  - Make sure you never drive voltage sources to fast and destroy your gates
Helpful
  - Automatically collect additional information about the data (“metadata”)
  - User should have to type in additional information only once and then never again

---
Lab::VISA is in Perl
Lab::VISA is in Perl
- Lab::VISA is written in Perl and C
- To be used from programs written in Perl
- Perl: interpreted scripting language
- Runs on almost every OS
- Extremely good in reading data files, manipulating data, etc.
- Allows to write quick and dirty scripts that get the job done
- Also allows to write clean fullsize programs
=> ideal for experimental physics

---
Lab::VISA architecture
Lab::VISA architecture
- Divided into three parts
- Build on top of each other
- Provide increasing comfort
- Measurement scripts can be based on any of these stages
Lab::VISA is the lowest layer. It makes the NI-VISA library accessible from perl and therefor allows to make any standard VISA call
Lab::Instrument package makes communication with instruments easier by silently handling the protocol involved
Lab::Tools is the highest abstraction layer. Provides support for writing good measurement scripts. Offers means of saving data and related meta information to disk, plotting data, etc.

---
Lab::VISA architecture
your (tiny) measurement program / script

---
Example 1
Example 1: Use only Lab::VISA
looks like example programs from the NI VISA manual
  - protocol overhead is pain in the neck. NOT GOOD.
#!/usr/bin/perl
use strict;
use Lab::VISA;
# Initialize VISA system and
# Open default resource manager
my ($status,$default_rm)=Lab::VISA::viOpenDefaultRM();
if ($status != $Lab::VISA::VI_SUCCESS) {
    die "Cannot open resource manager: $status";
}
# Open one resource (an instrument)
my $gpib=21; # we want to open the instrument
my $board=0; # with GPIB address 21
        # connected to GPIB board 0 in our computer
my $resource_name=sprintf("GPIB%u::%u::INSTR",$board,$gpib);
($status, my $instr)=Lab::VISA::viOpen(
    $default_rm, # the resource manager session
    $resource_name, # a string describing the
    $Lab::VISA::VI_NULL,# access mode (no special mode)
    $Lab::VISA::VI_NULL # time out for open (no time out)
);
if ($status != $Lab::VISA::VI_SUCCESS) {
    die "Cannot open instrument $resource_name. status: $status";
}
(...)
(...)
# Now we are going to send one command and read the result.
# We send the simple SCPI command "*IDN?" which asks the instrument
# to identify itself. Of course the instrument must support this
# command, in order to make this example work.
my $cmd="*IDN?";
($status, my $write_cnt)=Lab::VISA::viWrite(
    $instr, # the session identifier
    $cmd, # the command to send
    length($cmd) # the length of the command in bytes
);
if ($status != $Lab::VISA::VI_SUCCESS) {
    die "Error while writing: $status";
}
# Now we will read the instruments reply
($status, # indicates if the operation was successful
 my $result, # the answer string
 my $read_cnt)= # the length of the answer in bytes
    Lab::VISA::viRead(
       $instr, # the session identifier
       300 # read 300 bytes
    );
if ($status != $Lab::VISA::VI_SUCCESS) {
    die "Error while reading: $status";
}
# The result string will be 300 bytes long, but only $read_cnt
# bytes are part of the answer. We cut away the rest.
$result=substr($result,0,$read_cnt);
print $result;

---
Example 2
Example 2: Use Lab::Instrument
Much nicer!
#!/usr/bin/perl
use strict;
use Lab::Instrument::HP34401A;
my $gpib=21; # we want to open the instrument
my $board=0; # with GPIB address 21
        # connected to GPIB board 0 in our computer
# Create an instrument object
my $hp=new Lab::Instrument::HP34401A($board,$gpib);
# Use the id method to query the instruments ID string
my $result=$hp->id();
print $result;

---
Example 3
Example 3: My first real measurement!
#!/usr/bin/perl
use strict;
use Lab::Instrument::HP34401A;
use Lab::Instrument::Yokogawa7651;
# Connect to voltage source
my $yoko=new Lab::Instrument::Yokogawa7651({
    GPIB_address => 14,
    gate_protect => 0,
});
# Connect to multimeter
my $hp=new Lab::Instrument::HP34401A(0,21);
# Sweep Yokogawa from 0 to 1 volt (10 steps)
for (my $volt=0; $volt<=1; $volt=$volt+0.1) {
    # set Yokogawa
    $yoko->set_voltage($volt);
    # wait a second
    sleep(1);
    # read multimeter
    my $vmeas=$hp->read_voltage_dc();
    # print values
    print $volt,"\t",$vmeas,"\n";
}
- How about doing a gate sweep?
- That's easy!

---
Lab::Tools
Lab::Tools
- Provides additional tools to write better measurement scripts
- Store metadata alongside data
  - date and time
  - settings of additional instruments
  - ratio of voltage divider
  - color of the shirt you are wearing
  - everything that might be important for a later interpretation of the data
- Don‘t repeat yourself
  - Use above collected information automatically
  - Automatically plot data with correct axes, scaling, labels etc.

---
Metadata philosophy
Metadata philosophy
Constants
R = 1000 Ohm
Voltage source (Knick)
Vbias
Columns
C1: Vbias
C2: Vresistor
unit
description…
1.5	+8.19300500E-01
1.4	+7.23413000E-01
1.3	+6.28083900E-01
1.5	+8.19300500E-01
R
V
Multimeter (HP)
Vresistor
Axes
Axis “bias voltage”
(C1)
Axis “diode current”
(C2 / R)
Axis “diode resistance”
(R * (C1 / C2 – 1))
unit
expression
description…
Plots
Plot “diode current”
axis “diode current” over
axis “bias voltage”
Plot “diode resistance”
axis “diode resistance” over
axis “bias voltage”
logscale
grid
ranges…

---
Example 4
Example 4: Using Lab::Measurement
metadata, live plot
        {
        'unit' => 'Ohm',
        'expression' => 'R * ($C0/$C1 - 1)',
        'label' => 'R_{diode}',
        'description' => 'Diode resistance',
        },
    ],
    plots => {
        'diode current' => {
        'type' => 'line',
        'xaxis' => 0,
        'yaxis' => 1,
        'grid' => 'xtics ytics',
        },
        'diode resistance' => {
        'type' => 'line',
        'xaxis' => 0,
        'yaxis' => 2,
        'grid' => 'xtics ytics',
        'logscale' => 'y',
        },
    },
);
$measurement->start_block();
for (
    my $volt = $start_voltage;
    ($volt - $end_voltage) / $step < 0.5;
    $volt += $step
) {
    $knick->set_voltage($volt);
    usleep(500000);
    my $meas = $hp->read_voltage_dc(10,0.0001);
    $measurement->log_line($volt,$meas);
}
my $meta = $measurement->finish_measurement();
    constants => [
        {
        'name' => 'R',
        'value' => '1000',
        },
    ],
    columns => [
        {
        'unit' => 'V',
        'label' => 'V_{bias}',
        'description' => 'Bias Voltage',
        },
        {
        'unit' => 'V',
        'label' => 'Amplifier output',
        'description' => 'Voltage drop on serial resistor',
        }
    ],
    axes => [
        {
        'unit' => 'V',
        'expression' => '$C0',
        'label' => 'V_{bias}',
        'description' => 'Bias voltage',
        'min' => ($start_voltage < $end_voltage)
        ? $start_voltage
        : $end_voltage,
        'max' => ($start_voltage < $end_voltage)
        ? $end_voltage
        : $start_voltage,
        },
        {
        'unit' => 'mA',
        'expression' => '1000*($C1/R)',
        'label' => 'I_{diode}',
        'description' => 'Current through diode',
        'min' => '-1.2',
        'max' => '1.2',
        },
#!/usr/bin/perl
use strict;
use Lab::Instrument::KnickS252;
use Lab::Instrument::HP34401A;
use Lab::Measurement;
use Time::HiRes qw/usleep/;
################################
my $start_voltage = 1.5;
my $end_voltage = -3.5;
my $step = -0.05;
my $knick_gpib = 14;
my $hp_gpib = 21;
my $sample = "Zenerdiode";
my $title = "Messung mit Lab::Measurement";
my $comment = <<COMMENT;
Reihenschaltung aus Widerstand 1kOhm und Zenerdiode.
COMMENT
################################
my $knick=new Lab::Instrument::KnickS252({
    'GPIB_board' => 0,
    'GPIB_address' => $knick_gpib,
    'gate_protect' => 0,
});
my $hp=new Lab::Instrument::HP34401A(0,$hp_gpib);
my $measurement=new Lab::Measurement(
    sample => $sample,
    title => $title,
    filename_base => 'zener_kennlinie',
    description => $comment,
    live_plot => 'diode current',

---
More utilities: plotter.pl, make_overview.pl
- plotter.pl:
  - Reads measurement file, list avaible plots (axes etc)
  - Creates postscript or pdf output
- make_overview.pl:
  - Reads all measurements in a directory
  - Generates a postscript or pdf file with plots for each measurement, including the metadata (i.e. constants, parameters, color of your shirt) (via LaTeX)
  - Great for completing your lab book

---
Other Lab::VISA features
Other Lab::VISA features
“Gate protect” safety mechanism
  - Makes sure that no voltage is changed to fast
  - Big voltage steps are automatically split into small, slow steps
Date/time handling
  - Date/time column can contain timestamp for every data point
  - Plot data as function of time
- Measurements with higher dimensionality
  - Each trace/sweep/line is a “block”
  - Two-dimensional plots, selections of traces, etc.